## DMD
Test Group | Requirement (from Product Requirement & Trace Matrix) | Process/Steps | Hardware Required (or "N/A") | Desired Result | If Need FPGA Frimware function? | FPGA Firmware Requirement
8.1 Initial Power-up Test | 8.1.1 Program the FPGA with Test Firmware | 1) Connect the 10 pins header connector of the Actel FlashPro4 programming emulator to UUT J5. 2) Apply +24VDC between J18-1, 4 and J18-3, 6 3) Start programming 4) Verify FPGA programming passed | DC Power Supply Actel FlashPro4 programming emulator | FPGA programming passed | NO | N/A
| 8.1.2 24V DC Input | 1) Apply 24VDC 2.93A between J18-1, 4 and J18-3, 6(DC6# Source). Verified that LED D8 is on. 2) Reduce the voltage of DC#6 to 15V, Verified that LED D8 is off. 3) Reconfigure the voltage of DC#6 back to 24V | DC Power Supply FEASA | 1. DC Source output current within range 2. D8 will be on when the UUT is applied with 24VDC and be off when the UUT is applied with 15VDC | NO | N/A
| 8.1.3 Reset Circuit Test | 1) Verify the following LEDs D8 (24V), D30 (5V), D12 (-5VA), and D13 (+5VA) are ON. 2) Verify the LED D34 (POWER_GOOD) is ON. 3) Press and hold the switch SW1 4) Verify D34 (POWER_GOOD) is OFF. 5) Release the switch SW1. | DC Power Supply FEASA | 1. D8 (24V), D30 (5V), D12 (-5VA), and D13 (+5VA) will be on 2. D34 (POWER_GOOD) will be on when release the SW1 and be off when press the SW1 | NO | N/A
| 8.1.4 DSP to FPGA communication Test | 1) Write a test pattern 0xAA55 to FPGA’s scratchpad register. Perform a read back and verify that the received test pattern is 0xAA55. 2) Write a test pattern 0x55AA to FPGA’s scratchpad register. Perform a read back and verify that the received test pattern is 0x55AA. 3) Send powerIfRd Command to FPGA to verify that the status of signal 24V_GOOD# is low. | DC Power Supply | 1. Received test pattern is 0xAA55 when writing 0xAA55 to FPGA 2. Received test pattern is 0x55AA when writing 0x55AA to FPGA 3. 24V_GOOD# is low | Yes | Test patterns will be written to FPGA, FPGA will return result. Read the status of signal 24V_GOOD# and return result to PC.
| 8.1.5 Test Clock output of U55 and Watchdog U45 | 1) Send oscCounter Command to FPGA to verify the clock frequency of U55(FPGA_100M_CLK) to be 100MHz±0.1MHZ. 2) Send powerIfRd Command to FPGA to verify that the status of Watchdog output signal WDOK is high. | DC Power Supply | 1. The clock frequency of U55 is 100MHz±0.1MHZ 2. The status of Watchdog output signal WDOK is high(TP80) | Yes | Verify the clock frequency of U55(FPGA_100M_CLK) to be 100MHz±0.1MHZ and return result to PC.
| 8.1.6 Monitor UUT output voltages | 1) Monitor 12Viso output voltage via signal 12V_ISO_MON 2) Monitor +5VA output voltage via signal 5VA_POS_MON 3) Monitor -5VA output voltage via signal 5VA_NEG_MON 4) Monitor 24V output voltage via signal 24V_MON 5) Monitor 5V output voltage via signal 5V_MON 6) Monitor 5Viso output voltage via signal 5V_ISO_MON 7) Monitor 3.3V output voltage via signal 3.3V_MON 8) Monitor 2.5V output voltage via signal 2.5V_MON 9) Monitor 1.8V output voltage via signal 1.8V_MON 10) Monitor 1.0V output voltag via signal 1.0V_MON | DC Power Supply | 1) 12V_ISO_MON :12.0V±3% 2) 5VA_POS_MON :5V±3% 3) 5VA_NEG_MON : -5V±3%. 4) 24V_MON : 24V±5%. 5) 5V_MON : 5V±3% 6) 5V_ISO_MON : 5V±3% 7) 3.3V_MON :3.3V±2% 8) 2.5V_MON :2.5V±5% 9) 1.8V_MON :1.8V±5% 10) 1.0V_MON :1.0V±5% | Yes | FPGA control Analog Multiplexer to output certain voltage signal to ADC IC, monitor ADC output and return result to PC. 1) Monitor 12Viso output voltage via signal 12V_ISO_MON and return result to PC command 2) Monitor +5VA output voltage via signal 5VA_POS_MON and return result to PC command 3) Monitor -5VA output voltage via signal 5VA_NEG_MON and return result to PC command 4) Monitor 24V output voltage via signal 24V_MON and return result to PC command 5) Monitor 5V output voltage via signal 5V_MON and return result to PC command 6) Monitor 5Viso output voltage via signal 5V_ISO_MON and return result to PC command 7) Monitor 3.3V output voltage via signal 3.3V_MON and return result to PC command 8) Monitor 2.5V output voltage via signal 2.5V_MON and return result to PC command 9) Monitor 1.8V output voltage via signal 1.8V_MON and return result to PC command 10) Monitor 1.0V output voltag via signal 1.0V_MON and return result to PC command
| 8.1.8 Test Temperature Sensor Circuit U70 | Read the output voltage of the Temperature Sensor U70 at ambient temperature to be within 600mV to 960mV | DC Power Supply | U70 output(TP132) within 600mV to 960mV | Yes | FPGA control Analog Multiplexer to output certain voltage signal to ADC IC, monitor ADC output and return result to PC.
| 8.1.9 DAC_OUT signal Test at U62 Pin 5 | 1) Send dacWr command to FPGA to set DAC U62 output pin 5 to 0V. Send adcMux and adcSel to FPGA to select signal DAC_OUT to be the input of the ADC U43. Send adc Acquire command to FPGA to read the ADC output voltage to be between 0V±15mV. 2) Send dacWr command to FPGA to set DAC U62 output pin 5 to 5V. Send adcAcquire command to FPGA to read the ADC output voltage to be between 5V±5%. | DC Power Supply | 1) DAC U62 Pin 5 set to 0V, DAC_OUT set to be the input of the ADC U43, ADC output voltage(TP78) between 0V±15mV 2) DAC U62 Pin 5 set to 5V, ADC output voltage(TP78) between 5V±5% | Yes | 1) Set DAC U62 output pin 5 to 0V, select signal DAC_OUT to be the input of the ADC U43, read the ADC output voltage read the ADC output voltage and return result to PC. 2) Set DAC U62 output pin 5 to 5V, read the ADC output voltage read the ADC output voltage and return result to PC.
| 8.1.10 Miscellaneous Test | 1) Send gpiRd Command to FPGA to query the DMD Dash Number configuration to be 0x4. (To set the real expected value.) 2) Turns on DMD_24V power input(Apply +24VDC between J9-1, 3, 5 and J9-2, 4, 6). Send gpiRd Command to FPGA to query the status of signal SPD2_24V_STATUS# to be low. 3) Turns off DMD_24V power input. Send gpiRd Command to FPGA to query the status of signal SPD2_24V_STATUS# to be high. | DC Power Supply | 1) DMD Dash Number configuration is 0X4(To set the real expected value.) 2) SPD2_24V_STATUS# is be low when +24VDC between J9-1, 3, 5 and J9-2, 4, 6 applied 3) SPD2_24V_STATUS# is be low when +24VDC between J9-1, 3, 5 and J9-2, 4, 6 disconnected | Yes | 1) Query the DMD Dash Number configuration and return result to PC. 2) Query the status of signal SPD2_24V_STATUS# and return result to PC.
8.2 SPD DMD Interface Test | 8.2.1 | Connect J9.8, J9.10, J9.14 and J9.16 to J9.7, J9.9, J9.13 and J9.15 respectively | DC Power Supply | N/A | No | N/A
| 8.2.2 | Send gpoWr Command to FPGA to assert logic 1 to the following signals GNT_HW_EN_MON, LFT_HW_EN_MON, and DMD_PWR_OK. | DC Power Supply | GNT_HW_EN_MON, LFT_HW_EN_MON, and DMD_PWR_OK set to 0 | Yes | Assert logic 1 to the following signals GNT_HW_EN_MON, LFT_HW_EN_MON, and DMD_PWR_OK and return result to PC.
| 8.2.3 | Measure the voltage at J9.8 to be between 23.5V and 24.5V. | DC Power Supply DMM | Voltage at J9.8 is 24V±0.5V | NO | N/A
| 8.2.4 | Measure the voltage at J9.10 to be 0V±50mV. | DC Power Supply DMM | Voltage at J9.10 is 0V±50mV | NO | N/A
| 8.2.5 | Measure the voltage at J9.14 to be between 23.5V and 24.5V. | DC Power Supply DMM | Voltage at J9.14 is 24V±0.5V | NO | N/A
| 8.2.6 | Measure the voltage at J9.16 to be 0V±50mV. | DC Power Supply DMM | Voltage at J9.16 is 0V±50mV | NO | N/A
| 8.2.7 | Measure the voltage at J9.19 to be between 23.5V and 24.5V. | DC Power Supply DMM | Voltage at J9.19 is 24V±0.5V | NO | N/A
| 8.2.8 | Verify LEDs D81, D72, D71 and D70 are ON. | DC Power Supply FEASA | D81, D72, D71 and D70 are ON. | NO | N/A
| 8.2.9 | Send spdDmdRd command to FPGA to read and verify the status of signal SPDIO_FLT# is high. | DC Power Supply | SPDDIO_FLT# is high | Yes | Read the status of signal SPDIO_FLT# and return result to PC.
| 8.2.10 | Send spdDmdRd command to FPGA to read and verify the status of SPD DMD IF to be 0x14. | DC Power Supply | SPD DMD IF is 0x14 | Yes | Read the status of SPD DMD IF and return result to PC.
| 8.2.11 | Send gpoWr Command to FPGA to output logic 1 at signals LFT_ST_DISB_MON and GNT_ST_DISB_MON. | DC Power Supply | LFT_ST_DISB_MON and GNT_ST_DISB_MON set to logic 1 | Yes | Output logic 1 at signals LFT_ST_DISB_MON and GNT_ST_DISB_MON and return result to PC.
| 8.2.12 | Measure the voltage at J9.8 to be 0V±50mV. | DC Power Supply DMM | Voltage at J9.8 is 0V±50mV | NO | N/A
| 8.2.13 | Measure the voltage at J9.10 to be between 23.5V and 24.5V. | DC Power Supply DMM | Voltage at J9.10 is 24V±0.5V | NO | N/A
| 8.2.14 | Measure the voltage at J9.14 to be 0V± 50mV. | DC Power Supply DMM | Voltage at J9.14 is 0V±50mV | NO | N/A
| 8.2.15 | Measure the voltage at J9.16 to be between 23.5V and 24.5V. | DC Power Supply DMM | Voltage at J9.16 is 24V±0.5V | NO | N/A
| 8.2.16 | Measure the voltage at J9.19 to be 0V± 50mV. | DC Power Supply DMM | Voltage at J9.19 is 0V±50mV | NO | N/A
| 8.2.17 | Verify LEDs D81, D72, D71 and D70 are OFF. | DC Power Supply FEASA | D81, D72, D71 and D70 are OFF. | NO | N/A
| 8.2.18 | Send spdDmdRd command to FPGA to read and verify the status of SPD DMD IF to be 0xA. | DC Power Supply | SPD DMD IF is 0xA | Yes | Read the status of SPD DMD IF and return result to PC.
| 8.2.19 | Disconnect J9.8, J9.10, J9.14 and J9.16 from J9.7, J9.9, J9.13 and J9.15 respectively. | DC Power Supply | N/A | NO | N/A
| 8.2.20 | Send command gpoWr command to FPGA to output logic 1 to signal GNT_HW_EN_MON. | DC Power Supply | GNT_HW_EN_MON set to logic 1 | Yes | Output logic 1 to signal GNT_HW_EN_MON and return result to PC.
| 8.2.21 | Measure the voltage at J9.8 to be between 23.5V and 24.5V. | DC Power Supply DMM | Voltage at J9.8 is 24V±0.5V | NO | N/A
| 8.2.22 | Connect 30Ohms load to U30.20.(connect to J9.8) Measure the voltage at J9.8 to be between -500mV and 500mV. | DC Power Supply Resistive Load DMM | Voltage at J9.8 is 0V±500mV | NO | N/A
| 8.2.23 | Send spdDmdRd command to FPGA to verify the status of signal SPDIO_FLT# is low. | DC Power Supply Resistive Load | SPDDIO_FLT# is low | Yes | Read the status of signal SPDIO_FLT# and return result to PC.
| 8.2.24 | Disconnect 30 Ohms load from U30.20. | DC Power Supply | N/A | NO | N/A
| 8.2.25 | Send spdDmdRd command to FPGA to verify the status of signal SPDIO_FLT# is back to high. | DC Power Supply | SPDDIO_FLT# is high | Yes | Read the status of signal SPDIO_FLT# and return result to PC.
8.3 SPD EMOPS Interface Test | 8.3.1 | Connect J10.6, J10.14, J10.16, J10.18, and J10.19 to J10.1, J10.2, J10.24, DMD_24V and J10.25 respectively. | DC Power Supply | N/A | NO | N/A
| 8.3.2 | Send gpoWr command to FPGA to output logic 1 at signals EMPOS_STAT1, and LFT_ROT_BRK_RLS. Verify LEDs D16 and D22 are ON. | DC Power Supply FEASA | 1) EMPOS_STAT1, and LFT_ROT_BRK_RLS set to logic 1 2) D16 and D22 are ON | Yes | Output logic 1 at signals EMPOS_STAT1, and LFT_ROT_BRK_RLS and return result to PC.
| 8.3.3 | Measure the voltage at signal SPD_EMSTAT1 to be between 23.5V and 24.5V. (TP13) | DC Power Supply DMM | Voltage at SPD_EMSTAT1 is 24V±0.5V | NO | N/A
| 8.3.4 | Measure the voltage at signal SPD_EMSTAT2 to be 0V±50mV. (TP14) | DC Power Supply DMM | Voltage at SPD_EMSTAT2 is 0V±50mV | NO | N/A
| 8.3.5 | Measure the voltage at signal SPD_LFT_ROT_BRK_OVRD to be between 23.5V and 24.5V.(J10.24) | DC Power Supply DMM | Voltage at SPD_LFT_ROT_BRK_OVRD is 24V±0.5V | NO | N/A
| 8.3.6 | Measure the voltage at signal SPD_LAT_LNG_BRK_OVRD to be 0V±50mV. (J10.25) | DC Power Supply DMM | Voltage at SPD_LAT_LNG_BRK_OVRD is 0V±50mV | NO | N/A
| 8.3.7 | Send gpoWr command to FPGA to output logic 1 at signals EMOPS_STAT2, and LAT_LNG_BRK_RLS. Verify LEDs D17 and D24 are ON. | DC Power Supply FEASA | 1) EMOPS_STAT2, and LAT_LNG_BRK_RLS set to logic 1 2) D17 and D24 are ON | Yes | Output logic 1 at signals EMOPS_STAT2, and LAT_LNG_BRK_RLS and return result to PC.
| 8.3.8 | Measure the voltage at signal SPD_EMSTAT1 to be 0V±50mV. (TP13) | DC Power Supply DMM | Voltage at SPD_EMSTAT1 is 0V±50mV | NO | N/A
| 8.3.9 | Measure the voltage at signal SPD_EMSTAT2 to be between 23.5V and 24.5V. (TP14) | DC Power Supply DMM | Voltage at SPD_EMSTAT2 is 24V±0.5V | NO | N/A
| 8.3.10 | Measure the voltage at signal SPD_LFT_ROT_BRK_OVRD to be 0V±50mV.(J10.24) | DC Power Supply DMM | Voltage at SPD_LFT_ROT_BRK_OVRD is 0V±50mV | NO | N/A
| 8.3.11 | Measure the voltage at signal SPD_LAT_LNG_BRK_OVRD to be between 23.5V and 24.5V. (J10.25) | DC Power Supply DMM | Voltage at SPD_LAT_LNG_BRK_OVRD is 24V±0.5V | NO | N/A
| 8.3.12 | Send gpoWr command to FPGA to output logic 0 at signals EMOPS_STAT1, EMOPS_STAT2, LFT_ROT_BRK_RLS, and LAT_LNG_BRK_RLS. Verify LEDs D16, D17, D22 and D24 are OFF. | DC Power Supply FEASA | 1) EMOPS_STAT1, EMOPS_STAT2, LFT_ROT_BRK_RLS, and LAT_LNG_BRK_RLS set to 0 2) LEDs D16, D17, D22 and D24 are OFF. | Yes | Output logic 0 at signals EMOPS_STAT1, EMOPS_STAT2, LFT_ROT_BRK_RLS, and LAT_LNG_BRK_RLS and return result to PC.
| 8.3.13 | Measure the voltage at signals SPD_EMSTAT1, SPD_EMSTAT2, SPD_LFT_ROT_BRK_OVRD and SPD_LAT_BRK_OVRD to be 0V±50mV. | DC Power Supply DMM | Voltage at SPD_EMSTAT1, SPD_EMSTAT2, SPD_LFT_ROT_BRK_OVRD and SPD_LAT_BRK_OVRD is 0V±50mV. | NO | N/A
| 8.3.14 | Send gpoWr command to FPGA to output logic 1 at signals EMOPS_STAT2 and LAT_LNG_BRK_RLS. | DC Power Supply | EMOPS_STAT2 and LAT_LNG_BRK_RLS set to 1. | Yes | Output logic 1 at signals EMOPS_STAT2 and LAT_LNG_BRK_RLS and return result to PC.
| 8.3.15 | Send spdEmopsRd command to FPGA to read and verify the status of signals CCH_LFT_MOT_EN# and CCH_LAT_LNG_FLOAT# are low, while CCH_LFT_DWN# and EMO_GOOD# are high. | DC Power Supply | CCH_LFT_MOT_EN# and CCH_LAT_LNG_FLOAT# are low, while CCH_LFT_DWN# and EMO_GOOD# are high. | Yes | Read the status of signals CCH_LFT_MOT_EN#, CCH_LAT_LNG_FLOAT#, CCH_LFT_DWN# and EMO_GOOD# and return result to PC.
| 8.3.16 | Verify D42 and D44 are ON. | DC Power Supply FEASA | D42 and D44 are ON. | NO | N/A
| 8.3.17 | Send gpoWr command to FPGA to output logic 1 at signals EMOPS_STAT1 and LFT_ROT_BRK_RLS. | DC Power Supply | EMOPS_STAT1 and LFT_ROT_BRK_RLS set to 1. | Yes | Output logic 1 at signals EMOPS_STAT1 and LFT_ROT_BRK_RLS and return result to PC.
| 8.3.18 | Send spdEmopsRd command to FPGA to read and verify the status of signals CCH_LFT_MOT_EN# and CCH_LAT_LNG_FLOAT# are high, while CCH_LFT_DWN# and EMO_GOOD# are low. | DC Power Supply | CCH_LFT_MOT_EN# and CCH_LAT_LNG_FLOAT# are high, while CCH_LFT_DWN# and EMO_GOOD# are low. | Yes | Read the status of signals CCH_LFT_MOT_EN#,CCH_LAT_LNG_FLOAT# , CCH_LFT_DWN# and EMO_GOOD# and return result to PC.
| 8.3.19 | Verify D43 and D45 are ON. | DC Power Supply FEASA | D43 and D45 are ON. | NO | N/A
| 8.3.20 | Send spdEmopsRd command to FPGA to read and verify the status of signal SITE_24V_PWR_OK# is low. | DC Power Supply | SITE_24V_PWR_OK# is low. | Yes | Read the status of signal SITE_24V_PWR_OK# and return result to PC.
| 8.3.21 | Verify LED D46 is ON. | DC Power Supply FEASA | D46 is ON. | NO | N/A
| 8.3.22 | Disconnect J10.6, J10.14, J10.16, J10.18, and J10.19 from J10.1, J10.2, J10.24, DMD_24V and J10.25. | DC Power Supply | N/A | NO | N/A
| 8.3.23 | Send spdEmopsRd to FPGA to read and verify status of signal SITE_24V_PWR_OK# is high. | DC Power Supply | SITE_24V_PWR_OK# is high. | Yes | Read status of signal SITE_24V_PWR_OK# and return result to PC.
| 8.3.24 | Assert logic 1 to signal EM_CCH_24V_EN to enable U28. Measure the voltage at J10.26 (SPD_EM_24V_CCH) to be 24V±5%. | DC Power Supply DMM | 1) EM_CCH_24V_EN set to 1. 2) Voltage at J10.26 (SPD_EM_24V_CCH) is 24V±5%. | Yes | Assert logic 1 to signal EM_CCH_24V_EN to enable U28 and return result to PC.
| 8.3.25 | Assert logic 0 to signal EM_CCH_24V_EN to disable U28. Measure the voltage at J10.26 (SPD_EM_24V_CCH) to be 0V±0.5V. | DC Power Supply DMM | 1) EM_CCH_24V_EN set to 0. 2) Voltage at J10.26 (SPD_EM_24V_CCH) is 0V±0.5V. | Yes | Assert logic 0 to signal EM_CCH_24V_EN to enable U28 and return result to PC.
| 8.3.26 | Connect J10.5 to J10.1. Assert logic 1 to signal EMOPS_STAT1. Verify the status of signal EM_UP_LIMIT is low. | DC Power Supply | 1) EMOPS_STAT1 set to 1. 2) EM_UP_LIMIT is low. | Yes | Assert logic 1 to signal EMOPS_STAT1, Read the status of signal EM_UP_LIMIT and return result to PC.
| 8.3.27 | Verify D41 is on. | DC Power Supply FEASA | D41 is on. | NO | N/A
| 8.3.28 | Assert logic 0 to signal EMOPS_STAT1, verify the status of signal EM_UP_LIMIT is high. Disconnect J10.5 from J10.1. | DC Power Supply | 1) EMOPS_STAT1 set to 0. 2) EM_UP_LIMIT is high. | Yes | Assert logic 0 to signal EMOPS_STAT1, Read the status of signal EM_UP_LIMIT and return result to PC.
| 8.3.29 | Connect J10.8 to J10.1. Assert logic 1 to signal EMOPS_STAT1, verify the status of signal EM_DOWN_LIMIT is low. | DC Power Supply | 1) EMOPS_STAT1 set to 1. 2) EM_DOWN_LIMIT is low. | Yes | Assert logic 1 to signal EMOPS_STAT1, Read the status of signal EM_DOWN_LIMIT and return result to PC.
| 8.3.30 | Verify D35 is on. | DC Power Supply FEASA | D35 is on. | NO | N/A
| 8.3.31 | Assert logic 0 to signal EMOPS_STAT1, verify the status of signal EM_DOWN_LIMIT is high. Disconnect J10.8 from J10.1. | DC Power Supply | 1) EMOPS_STAT1 set to 0. 2) EM_DOWN_LIMIT is high. | Yes | Assert logic 0 to signal EMOPS_STAT1, Read the status of signal EM_DOWN_LIMIT and return result to PC.
| 8.3.32 | Send command gpoWr command to FPGA to output logic 1 to signal EMOPS_STAT1. | DC Power Supply | EMOPS_STAT1 set to 1. | Yes | Output logic 1 to signal EMOPS_STAT1 and return result to PC.
| 8.3.33 | Measure the voltage at J10.1 to be between 23.5V and 24.5V. | DC Power Supply DMM | Voltage at J10.1 is 24V±0.5V. | NO | N/A
| 8.3.34 | Connect 30Ohms load to U27.20(TP13/J10.1). Measure the voltage at J10.1 to be between -50mV and 50mV. | DC Power Supply Resistive Load DMM | Voltage at J10.1 is 0V±50mV. | NO | N/A
| 8.3.35 | Send spdDmdRd command to FPGA to verify the status of signal SPD_EMOPS_FLT# is low. | DC Power Supply Resistive Load | SPD_EMOPS_FLT# is low. | Yes | Read the status of signal SPD_EMOPS_FLT# and return result to PC.
| 8.3.36 | Disconnect 30 Ohms load from U27.20. | DC Power Supply | N/A | NO | N/A
| 8.3.37 | Send spdDmdRd command to FPGA to verify the status of signal SPD_EMOPS_FLT# is back to high. | DC Power Supply | SPD_EMOPS_FLT# is high. | Yes | Read the status of signal SPD_EMOPS_FLT# and return result to PC.
8.4 Stand Controller Interface Test | 8.4.1 | Send gpoWr command to FPGA to apply logic 1 at signal EXOPS_GNT_24V_EN to enable U28. Measure and verify the output voltage of U28(TP40) at J1.1 to be 24V±5%. | DC Power Supply DMM | Voltage at U28(TP40) J1.1 is 24V±5%. | Yes | Apply logic 1 at signal EXOPS_GNT_24V_EN to enable U28 and return result to PC.
| 8.4.2 | Send gpoWr command to FPGA to apply logic 0 at signal EXOPS_GNT_24V_EN to disable U28. Measure and verify the output voltage of U28 at J1.1 to be 0V±0.5V. | DC Power Supply DMM | Voltage at U28(TP40) J1.1 is 0V±0.5V. | Yes | Apply logic 0 at signal EXOPS_GNT_24V_EN to disable U28 and return result to PC.
| 8.4.3 | Apply +24V to J1.24 (CW_LIMIT_24V#) and 0V to J1.22 (CCW_LIMIT_24V#). | DC Power Supply | N/A | NO | N/A
| 8.4.4 | Send standContRd command to FPGA to verify the status of signal GANT_CW_LIMIT# is low, while signals GANT_CCW_LIMIT# is high. | DC Power Supply | 1) GANT_CW_LIMIT# is low, while signals 2) GANT_CCW_LIMIT# is high. | Yes | Read the status of signal GANT_CW_LIMIT#, GANT_CCW_LIMIT# and return result to PC.
| 8.4.5 | Verify D82 is OFF while D83 is ON. | DC Power Supply FEASA | D82 is OFF while D83 is ON. | NO | N/A
| 8.4.6 | Apply 0V to J1.24 (CW_LIMIT_24V#) and +24V to J1.22 (CCW_LIMIT_24V#). | DC Power Supply | N/A | NO | N/A
| 8.4.7 | Send standContRd command to FPGA to verify the status of signal GANT_CW_LIMIT# is high, while signals GANT_CCW_LIMIT# is low. | DC Power Supply | 1) GANT_CW_LIMIT# is high, while signals 2) GANT_CCW_LIMIT# is low. | Yes | Read the status of signal GANT_CW_LIMIT#, signals GANT_CCW_LIMIT# and return result to PC.
| 8.4.8 | Verify D82 is ON while D83 is OFF. | DC Power Supply FEASA | D82 is ON while D83 is OFF | NO | N/A
| 8.4.9 | Remove voltage from J1.24 (CW_LIMIT_24V#) and J1.22 (CCW_LIMIT_24V#). | DC Power Supply | N/A | NO | N/A
| 8.4.10 | Loop J1.10 (MSSB_TX_P), J1.11(MSSB_TX_N) and J1.14(MSSB_RX_P), J1.15 (MSSB_RX_N). | DC Power Supply | N/A | NO | N/A
| 8.4.11 | Send command to make FPGA transits the packets, transmission rate is around 500K packets per second, with continuous back-to-back TX/RX. | DC Power Supply | Loop-back | Yes | Transits the packets, transmission rate is around 500K packets per second, with continuous back-to-back TX/RX. Perform loop-back test where connect the TX output to RX input and check FPGA registers. FPGA continuously checks for communication errors - transmission rate is around 500K packets per second, with continuous back-to-back TX/RX.
| 8.4.12 | Send command to check FPGA registers bits for Good Packets, Bad Packets and Aligned. Good Packets expect continuous increment, Bad Packets expect remain static and Aligned expect ‘1’. | DC Power Supply | Good Packets should continuously increment (spinning). Bad Packets should remain static Aligned bit should be '1' | Yes | There are three registers to check at address 24 (for stand-controller) and 25 (for service panel). Good packets (bits 15..0) Bad Packets (bits 27..16) Aligned (bit 31) Criteria: Good Packets should continuously increment (spinning). Bad Packets should remain static Aligned bit should be '1' [MSSB communication test section needs the test firmware function support, the command used to transit packets and check register bits is To Be Decided.]
8.6 CCHL Interface Test | 8.6.1 | Send motgpoWr command to FPGA to output logic 1 to signals LFT_SER_CLK, LFT_SER_SYNC, LFT_SER_DATA [1:0]. | DC Power Supply | LFT_SER_CLK, LFT_SER_SYNC, LFT_SER_DATA [1:0] set to 1 | Yes | Output logic 1 to signals LFT_SER_CLK, LFT_SER_SYNC, LFT_SER_DATA [1:0] and return result to PC.
| 8.6.2 | Connect J12.1, J12.2, J12.3, and J12.4 to J12.6, J12.10, J12.11, and J12.14 respectively. | DC Power Supply | N/A | NO | N/A
| 8.6.3 | Measure the voltage at signals LFT_GPI [4:1](J12.1, J12.2, J12.3, and J12.4?) to be within 23.5V to 24.5V. | DC Power Supply DMM | 1) Voltage at J12.1 is 24V±0.5V. 2) Voltage at J12.2 is 24V±0.5V. 3) Voltage at J12.3 is 24V±0.5V. 4) Voltage at J12.4 is 24V±0.5V. | NO | N/A
| 8.6.4 | Verify LEDs D60 and D61are ON. | DC Power Supply FEASA | D60 and D61are ON. | NO | N/A
| 8.6.5 | Send cchlRd command to FPGA to read and verify the status of signal LFT_SERIO_FLT# is high. | DC Power Supply | LFT_SERIO_FLT# is high. | Yes | Read the status of signal LFT_SERIO_FLT# and return result to PC.
| 8.6.6 | Send motgpoWr command to FPGA to output logic 1 to signals LFT_SER_CLK, LFT_SER_DATA0, and logic 0 to signals LFT_SER_SYNC, LFT_SER_DATA1. | DC Power Supply | 1) LFT_SER_CLK, LFT_SER_DATA0 set to 1. 2) LFT_SER_SYNC, LFT_SER_DATA1 set to 0. | Yes | Output logic 1 to signals LFT_SER_CLK, LFT_SER_DATA0, and logic 0 to signals LFT_SER_SYNC, LFT_SER_DATA1 and return result to PC.
| 8.6.7 | Send cchlRd command to FPGA to read and verify the status of CHHL_IF to be 0x5. | DC Power Supply | CHHL_IF is 0x5. | Yes | Read the status of CHHL_IF and return result to PC.
| 8.6.8 | Send motgpoWr command to FPGA to output logic 1 to signals LFT_SER_SNYC, LFTSER_DATA1 and logic 0 to LFT_SER_CLK, LFT_SER_DATA0. | DC Power Supply | 1) LFT_SER_CLK, LFT_SER_DATA0 set to 0. 2) LFT_SER_SYNC, LFT_SER_DATA1 set to 1. | Yes | Output logic 1 to signals LFT_SER_SNYC, LFTSER_DATA1 and logic 0 to LFT_SER_CLK, LFT_SER_DATA0 and return result to PC.
| 8.6.9 | Send cchlRd command to FPGA to read and verify the status of CCHL_IF to be 0xA. | DC Power Supply | CCHL_IF is 0xA. | Yes | Read the status of CCHL_IF and return result to PC.
| 8.6.10 | Send motgpoWr command to FPGA to output logic 1 signals LFT_SER_CLK, LFT_SER_DATA0, and logic 0 to signals LFT_SER_SYNC, LFT_SER_DATA1. Measure the voltage at signal LIFT_GPI1 to be within 23.5V and 24.5V. (J12.1) | DC Power Supply DMM | 1) LFT_SER_CLK, LFT_SER_DATA0 set to 1. 2) LFT_SER_SYNC, LFT_SER_DATA1 set to 0. 3) Voltage at LIFT_GPI1(J12.1) is 24V±0.5V. | Yes | Output logic 1 to signals LFT_SER_CLK, LFT_SER_DATA0, and logic 0 to signals LFT_SER_SYNC, LFT_SER_DATA1 and return result to PC.
| 8.6.11 | Connect 30Ohms resistive load to J12.1. Measure the voltage at LIFT_GPI1 to be 0V±50mV. | DC Power Supply Resistive Load DMM | Voltage at LIFT_GPI1(J12.1) is 0V±50mV. | NO | N/A
| 8.6.12 | Send cchlRd command to FPGA to read and verify the status of signal LFT_SERIO_FLT# to be low. | DC Power Supply Resistive Load | LFT_SERIO_FLT# is low. | Yes | Read the status of signal LFT_SERIO_FLT# and return result to PC.
| 8.6.13 | Disconnect 30Ohms resistive load from J12.1. | DC Power Supply | N/A | NO | N/A
| 8.6.14 | Send cchlRd command to FPGA to read and verify the status of signal LFT_SERIO_FLT# is back to high. | DC Power Supply | LFT_SERIO_FLT# is back to high. | Yes | Read the status of signal LFT_SERIO_FLT# and return result to PC.
| 8.6.15 | Disconnect J12.1, J12.2, J12.3, and J12.4 from J12.6, J12.10, J12.11, and J12.14. | DC Power Supply | N/A | NO | N/A
| 8.6.16 | Connect 100KOhms resistive between J12.23 and J12.24. | DC Power Supply Resistive Load | N/A | NO | N/A
| 8.6.17 | Send dacWr command to FPGA to set the output of LFT_DAC to 0V. Measure the voltage at signal LIFT_CURR_FB_A to be 0V±50mV. (J12.23) | DC Power Supply Resistive Load DMM | 1) LFT_DAC output set to 0V 2) Voltage at LIFT_CURR_FB_A (J12.23) is 0V±50mV. | Yes | Set the output of LFT_DAC to 0V and return result to PC.
| 8.6.18 | Set the output voltage of LFT_DAC to 2.5V. Measure the voltage at signal LIFT_CURR_FB_A to be 2.5V±4%. | DC Power Supply Resistive Load DMM | 1) LFT_DAC output set to 2.5V 2) Voltage at LIFT_CURR_FB_A (J12.23) is 2.5V±4%. | Yes | Set the output of LFT_DAC to 2.5V and return result to PC.
| 8.6.19 | Set the output voltage of LFT_DAC to 5.0V. Measure the voltage at signal LIFT_CURR_FB_A to be 5V±2%. | DC Power Supply Resistive Load DMM | 1) LFT_DAC output set to 5V 2) Voltage at LIFT_CURR_FB_A (J12.23) is 5V±2%. | Yes | Set the output of LFT_DAC to 5V and return result to PC.
| 8.6.20 | Set the output voltage of LFT_DAC to 0V. Disconnect 100KOhms resistive load from J12.23 and J12.24. | DC Power Supply DMM | LFT_DAC output set to 2.5V | Yes | Set the output of LFT_DAC to 0V and return result to PC.
| 8.6.21 | Apply 0V analog input between J12.21and J12.22. Select signal LFT_CURR_CMD to be the input of ADC. Send adcAquire command to FPGA to query and verify the voltage at signal LFT_CURR_CMD to be 0V±50mV. | DC Power Supply PXI-6704 | Voltage at LFT_CURR_CMD is 0V±50mV. | Yes | Select signal LFT_CURR_CMD to be the input of ADC, query and verify the voltage at signal LFT_CURR_CMD to be 0V±50mV and return result to PC.
| 8.6.22 | Set the analog input voltage to 2.5V. Send adcAquire command to FPGA to query and verify the voltage at signal LFT_CURR_CMD to be 0.625V±20mV. | DC Power Supply PXI-6704 | Voltage at LFT_CURR_CMD is 0.625V±20mV. | Yes | Select signal LFT_CURR_CMD to be the input of ADC, query the voltage at signal LFT_CURR_CMD and return result to PC.
| 8.6.23 | Set the analog input voltage to 5.0V. Send adcAquire command to FPGA to query and verify the voltage at signal LFT_CURR_CMD to be 1.25V±20mV. | DC Power Supply PXI-6704 | Voltage at LFT_CURR_CMD is 1.25V±20mV. | Yes | Select signal LFT_CURR_CMD to be the input of ADC, query the voltage at signal LFT_CURR_CMD and return result to PC.
| 8.6.24 | Set the analog input voltage to -5V. Send adcAquire command to FPGA to query and verify the voltage at signal LFT_CURR_CMD to be -1.25V±20mV. | DC Power Supply PXI-6704 | Voltage at LFT_CURR_CMD is -1.25V±20mV. | Yes | Select signal LFT_CURR_CMD to be the input of ADC, query the voltage at signal LFT_CURR_CMD and return result to PC.
| 8.6.25 | Set the analog input voltage to -2.5V. Send adcAquire command to FPGA to query and verify the voltage at signal LFT_CURR_CMD to be -0.625V±20mV. | DC Power Supply PXI-6704 | Voltage at LFT_CURR_CMD is -0.625V±20mV. | Yes | Select signal LFT_CURR_CMD to be the input of ADC, query the voltage at signal LFT_CURR_CMD and return result to PC.
| 8.6.26 | Set the analog input voltage to 0V. Send adcAquire command to FPGA to query and verify the voltage at signal LFT_CURR_CMD to be 0V±50mV. | DC Power Supply PXI-6704 | Voltage at LFT_CURR_CMD is 0V±50mV. | Yes | Select signal LFT_CURR_CMD to be the input of ADC, query the voltage at signal LFT_CURR_CMD and return result to PC.
8.7 Lift Motor Hall Sensor IF Test | 8.7.1 | Apply 20 Ohms resistive load between HALL_LFT_+5V and HALL_LFT_GND. | DC Power Supply Resistive Load | N/A | NO | N/A
| 8.7.2 | Send gpoWr command to FPGA to output logic 1 to signal LFT_HALL_PWR_EN#. | DC Power Supply Resistive Load | LFT_HALL_PWR_EN# set to 1 | Yes | Output logic 1 to signal LFT_HALL_PWR_EN# and return result to PC.
| 8.7.3 | Measure the voltage at HALL_LFT_+5V to be 0V±50mV. (J6.5) | DC Power Supply Resistive Load DMM | Voltage at HALL_LFT_+5V is 0V±50mV. | NO | N/A
| 8.7.4 | Send gpoWr command to FPGA to output logic 0 to signal LFT_HALL_PWR_EN#. Measure the voltage at signal HALL_LFT_+5V to be 5V±2%. | DC Power Supply Resistive Load DMM | 1) LFT_HALL_PWR_EN# set to 0 2) Voltage at HALL_LFT_+5V is 5V±2%. | Yes | Output logic 0 to signal LFT_HALL_PWR_EN# and return result to PC.
| 8.7.5 | Send liftMotRd command to FPGA to read and verify the status of signal LFT_HALL_PWR_OK is high. | DC Power Supply Resistive Load | LFT_HALL_PWR_OK is high. | Yes | Read the status of signal LFT_HALL_PWR_OK and return result to PC.
| 8.7.6 | Replace the 20 Ohms resistive load with 2 Ohms resistive load. | DC Power Supply Resistive Load | N/A | NO | N/A
| 8.7.7 | Measure the voltage at HALL_LFT_+5V to be 0V±50mV. | DC Power Supply Resistive Load DMM | Voltage at HALL_LFT_+5V is 0V±50mV. | NO | N/A
| 8.7.8 | Send liftMotRd command to FPGA to read and verify the status of signal LFT_HALL_PWR_OK is low. | DC Power Supply Resistive Load | LFT_HALL_PWR_OK is low. | Yes | Read the status of signal LFT_HALL_PWR_OK and return result to PC.
| 8.7.9 | Disconnect 2 Ohms resistive load from J6.5. Send gpoWr command to FPGA to output logic 1 to signal LFT_HALL_PWR_EN#. | DC Power Supply | LFT_HALL_PWR_EN# set to 1 | Yes | Output logic 1 to signal LFT_HALL_PWR_EN# and return result to PC.
| 8.7.10 | Connect 20 Ohms resistive load across HALL_LFT_+5V and HALL_LFT_GND. | DC Power Supply Resistive Load | N/A | NO | N/A
| 8.7.11 | Send gpoWr command to FPGA to output logic 0 to LFT_HALL_PWR_EN#. | DC Power Supply Resistive Load | LFT_HALL_PWR_EN# set to 0 | Yes | Output logic 0 to signal LFT_HALL_PWR_EN# and return result to PC.
| 8.7.12 | Measure the voltage at HALL_LFT_+5V to be 5V±10%. | DC Power Supply DMM | Voltage at HALL_LFT_+5V is 5V±10%. | NO | N/A
| 8.7.13 | Send liftMotRd command to FPGA to read and verify the status of signal LFT_HALL_PWR_OK is high. | DC Power Supply Resistive Load | LFT_HALL_PWR_OK is high. | Yes | Read the status of signal LFT_HALL_PWR_OK and return result to PC.
| 8.7.14 | Send gpoWr command to FPGA to output logic 1 to signal LFT_HALL_PWR_EN#. | DC Power Supply Resistive Load | LFT_HALL_PWR_EN# set to 1 | Yes | Output logic 1 to signal LFT_HALL_PWR_EN# and return result to PC.
| 8.7.15 | Remove 20 Ohms resistive load between HALL_LFT_+5V and HALL_LFT_GND. (J6.1) | DC Power Supply | N/A | NO | N/A
| 8.7.16 | Apply DIO logic 1 to signals IN110, and IN112. | DC Power Supply PXI-6704 | 1) IN110 set to 1. 2) IN112 set to 1. | NO | N/A
| 8.7.17 | Send liftMotRd command to FPGA to read and verify the status of signals LFT_HALL_SNS [3:1] to be 0x2. | DC Power Supply | LFT_HALL_SNS [3:1] is 0x2. | Yes | Read the status of signals LFT_HALL_SNS [3:1] and return result to PC.
| 8.7.18 | Verify LED D14 and D6 are ON. | DC Power Supply FEASA | D14 and D6 are ON. | NO | N/A
| 8.7.19 | Apply DIO logic 1 to signal IN111 | DC Power Supply PXI-6704 | IN111 set to 1. | NO | N/A
| 8.7.20 | Send liftMotRd command to FPGA to read and verify the status of signals LFT_HALL_SNS# [3:1] to be 0x5 | DC Power Supply | LFT_HALL_SNS [3:1] is 0x5. | Yes | Read the status of signals LFT_HALL_SNS [3:1] and return result to PC.
| 8.7.21 | Verify LED D7 is ON. Set the digital input at signal IN111 to logic 0. | DC Power Supply FEASA | 1) D7 is ON. 2) IN111 set to 0. | NO | N/A
8.8 Service Pendant Test | 8.8.1 | Send gpoWr command to FPGA to apply logic 1 at signal PAN_24V_SW to enable U28. Measure and verify the output voltage of U28 at J8.1 to be 24V±5%. | DC Power Supply DMM | 1) PAN_24V_SW set to 1. 2) Voltage at J8.1 is 24V±5%. | Yes | Apply logic 1 at signal PAN_24V_SW to enable U28 and return result to PC.
| 8.8.2 | Send gpoWr command to FPGA to apply logic 0 at signal PAN_24V_SW to disable U28. Measure and verify the output voltage of U28 at J8.1 to be 0V±0.5V. | DC Power Supply DMM | 1) PAN_24V_SW set to 0. 2) Voltage at J8.1 is 0V±0.5V. | Yes | Apply logic 0 at signal PAN_24V_SW to enable U28 and return result to PC.
| 8.8.3 | Loop J8.5 (SRV_MSSB_RS422_TX+), J8.10 (SRV_MSSB_RS422_TX-) and J8.4 (SRV_MSSB_RS422_RX+), J8.9 (SRV_MSSB_RS422_RX-). | DC Power Supply | N/A | NO | N/A
| 8.8.4 | Send command to make FPGA transits the packets, transmission rate is around 500K packets per second, with continuous back-to-back TX/RX. | DC Power Supply | Loop-back | Yes | Transits the packets, transmission rate is around 500K packets per second, with continuous back-to-back TX/RX. Perform loop-back test where connect the TX output to RX input and check FPGA registers. FPGA continuously checks for communication errors - transmission rate is around 500K packets per second, with continuous back-to-back TX/RX.
| 8.8.5 | Send command to check FPGA registers bits for Good Packets, Bad Packets and Aligned. Good Packets expect continuous increment, Bad Packets expect remain static and Aligned expect ‘1’. | DC Power Supply | Good Packets should continuously increment (spinning). Bad Packets should remain static Aligned bit should be '1' | Yes | There are three registers to check at address 24 (for stand-controller) and 25 (for service panel). Good packets (bits 15..0) Bad Packets (bits 27..16) Aligned (bit 31) Criteria: Good Packets should continuously increment (spinning). Bad Packets should remain static Aligned bit should be '1' [MSSB communication test section needs the test firmware function support, the command used to transit packets and check register bits is To Be Decided.]
8.9 Gantry Motor Test | 8.9.1 | Apply 24V to EMOS_24V(J18?). Connect J14.4 (GNT_96VSHUNT-) to scope CH3, connect 25ohms resistor across J14.3 and J14.4. | DC Power Supply Oscilloscope Resistive Load | N/A | NO | N/A
| 8.9.2 | Apply 96VDC between signal 96V_GNT(J17.2,J17.3) and MGND | DC Power Supply Resistive Load | N/A | NO | N/A
| 8.9.3 | Send gantry96VoltsEn command to FPGA to output logic 1 to signal 24V_GNT_EMOPS_EN and GNT_MOT_PWR_EN, measure the voltage at J14.3 to be 24V±5%. | DC Power Supply DMM Resistive Load | 1) 24V_GNT_EMOPS_EN and GNT_MOT_PWR_EN set to 1. 2) Voltage at J14.3 is 24V±5%. | Yes | Output logic 1 to signal 24V_GNT_EMOPS_EN and GNT_MOT_PWR_EN and return result to PC.
| 8.9.4 | Send gantry96VoltsEn command to FPGA to output logic 0 to signal 24V_GNT_EMOPS_EN and logic 1 to signal GNT_MOT_PWR_EN. Measure the voltage at J14.3 and GND to be 96V±5%. | DC Power Supply DMM Resistive Load | 1) 24V_GNT_EMOPS_EN set to 0 and GNT_MOT_PWR_EN set to 1. 2) Voltage at J14.3 is 96V±5%. | Yes | Output logic 0 to signal 24V_GNT_EMOPS_EN and logic 1 to signal GNT_MOT_PWR_EN and return result to PC.
| 8.9.5 | Send gantry96VoltsEn command to FPGA to output logic 0 to signal 24V_GNT_EMOPS_EN and GNT_MOT_PWR_EN, measure the voltage at J14.3 to be 0Vdc to 0.5Vdc. | DC Power Supply DMM Resistive Load | 1) 24V_GNT_EMOPS_EN and GNT_MOT_PWR_EN set to 0 2) Voltage at J14.3 is 0V to 0.5V. | Yes | Output logic 0 to signal 24V_GNT_EMOPS_EN and GNT_MOT_PWR_EN and return result to PC.
| 8.9.6 | Send gantry96VoltsEn command to FPGA to output logic 1 to signal GNT_MOT_PWR_EN. Measure the voltage at J14.3 and GND to be 96V±5%. | DC Power Supply DMM Resistive Load | 1) GNT_MOT_PWR_EN set to 1. 2) Voltage at J14.3 is 96V±5%. | Yes | Output logic 1 to signal GNT_MOT_PWR_EN and return result to PC.
| 8.9.7 | Send command to FPGA to output logic 1 to GNT_SHUNT_EN_OUT for 2.5s | DC Power Supply Resistive Load | GNT_SHUNT_EN_OUT set to 1 for 2.5s. | Yes | Output logic 1 to signal GNT_SHUNT_EN_OUT for 2.5s and return result to PC.
| 8.9.8 | Measure the voltage at J14.4 (GNT_96VSHUNT-) from Scope CH3, it should be a 96V-0V pulse for 2.5s period. | DC Power Supply Oscilloscope Resistive Load | Voltage at J14.4 is a 96V-0V pulse for 2.5s period. | NO | N/A
| 8.9.9 | Disconnect 25ohms resistor from J14.3, disconnect CH3 from J14.4. | DC Power Supply | N/A | NO | N/A
| 8.9.10 | Apply 24VDC to signal EMOPS_24V and GND. Send powerIfRd command to FPGA to read and verify the status of signal 24V_GNT_EMOPS_PG is low | DC Power Supply | 24V_GNT_EMOPS_PG is low. | Yes | Read the status of signals signal 24V_GNT_EMOPS_PG and return result to PC.
| 8.9.11 | Send gantry96VoltsEn to FPGA to output logic 1 to 24V_GNT_EMOPS_EN. | DC Power Supply | 24V_GNT_EMOPS_EN set to 1 | Yes | Output logic 1 to signal 24V_GNT_EMOPS_EN and return result to PC.
| 8.9.12 | Send powerIfRd command to FPGA to read and verify the status of signal 24V_GNT_EMOPS_PG is high. | DC Power Supply | 24V_GNT_EMOPS_PG is high. | Yes | Read the status of signals signal 24V_GNT_EMOPS_PG and return result to PC.
| 8.9.13 | Select signal 24V_IN_MON to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage at signal 24V_IN to be 24V±5%. | DC Power Supply | 1) 24V_IN_MON is the input of ADC. 2) Voltage at 24V_IN is 24V±5%. | Yes | Select signal 24V_IN_MON to be the input of ADC, query the voltage at signal 24V_IN and return result to PC.
| 8.9.14 | Select signal 24V_GNT_EMOPS_S_MON to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage at signal 24V_GNT_EMOPS_S to be 24V±5%. | DC Power Supply | 1) 24V_GNT_EMOPS_S_MON is the input of ADC. 2) Voltage at 24V_GNT_EMOPS_S is 24V±5%. | Yes | Select signal 24V_GNT_EMOPS_S_MON to be the input of ADC, query the voltage at signal 24V_GNT_EMOPS_S and return result to PC.
| 8.9.15 | Select signal 96V_24V_GNT_MON to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage at signal 96V_24V_GNT 21.5V and 24.5V. | DC Power Supply | 1) 96V_24V_GNT_MON is the input of ADC. 2) Voltage at 96V_24V_GNT is 23V±1.5V. | Yes | Select signal 96V_24V_GNT_MON to be the input of ADC, query the voltage at signal 96V_24V_GNT and return result to PC.
| 8.9.16 | Verify the output voltage of U99 is between 0.6V and 0.96V | DC Power Supply | Voltage at GNT_DRVR_TEMP is 0.6V ~ 0.96V. | Yes | Query the voltage at signal GNT_DRVR_TEMP and return result to PC.
| 8.9.17 | Send gantMotPwmSel and gantMotEn command to PFGA to output 333.3Hz pulses to signals GNT_PWM_PHA_HI/LO, GNT_PWM_PHB_HI/LO and GNT_PWM_PHC_HI/LO. | DC Power Supply | GNT_PWM_PHA_HI/LO, GNT_PWM_PHB_HI/LO and GNT_PWM_PHC_HI/LO get 333.3Hz pulses | Yes | Output 333.3Hz pulses to signals GNT_PWM_PHA_HI/LO, GNT_PWM_PHB_HI/LO and GNT_PWM_PHC_HI/LO and return result to PC.
| 8.9.18 | Send gantMotGpio command to FPGA to output logic 1 at signal GNT_PWM_EN. | DC Power Supply | GNT_PWM_EN set to 1. | Yes | Output logic 1 at signal GNT_PWM_EN and return result to PC.
| 8.9.19 | Measure the signal amplitude of GNT_DRV_PHASE_A, B, and C to be 24V±10%. | DC Power Supply Oscilloscope | Amplitude of GNT_DRV_PHASE_A, B, and C is 24V±10%. | NO | N/A
| 8.9.20 | Measure the signal frequency of GNT_DRV_PHASE_A, B, and C to be 333.3Hz±5%. | DC Power Supply Oscilloscope | Frequency of GNT_DRV_PHASE_A, B, and C is 333.3Hz±5%. | NO | N/A
| 8.9.21 | Measure the phase between GNT_DRV_PHASE_A, B, and C to be within 115 and 125 degrees. | DC Power Supply Oscilloscope | Phase between GNT_DRV_PHASE_A, B, and C is within 115 and 125 degrees. | NO | N/A
| 8.9.22 | Send gantMotGpio command to FPGA to output logic 0 to signal GNT_PWM_EN. | DC Power Supply | GNT_PWM_EN set to 0. | Yes | Output logic 0 at signal GNT_PWM_EN and return result to PC.
| 8.9.23 | Connect 25Ohms resistive loads delta connection to GNT_DRV_PHASE_A, B, and C. | DC Power Supply Resistive Load | N/A | NO | N/A
| 8.9.24 | Send gantMotGpio command to FPGA to output logic 1 to signal GNT_PWM_EN. | DC Power Supply Resistive Load | GNT_PWM_EN set to 1. | Yes | Output logic 1 at signal GNT_PWM_EN and return result to PC.
| 8.9.25 | Select the signal 96V_24V_GNT_IMON_V to be an input of the ADC. Send adcAcquire command to FPGA to query and verify the voltage at signal 96V_24V_GNT_IMON_V to be within 87.5mV and 112.5mV. (Spec limit to be settled with KGB.) | DC Power Supply Resistive Load | Voltage at 96V_24V_GNT_IMON_V is 87.5mV ~ 112.5mV. | Yes | Select the signal 96V_24V_GNT_IMON_V to be an input of the ADC, query the voltage at signal 96V_24V_GNT_IMON_V and return result to PC.
| 8.9.26 | Send gantry96VoltsEn command to FPGA to output logic 0 to signal GNT_MOT_PWR_EN. | DC Power Supply Resistive Load | GNT_MOT_PWR_EN set to 0. | Yes | Output logic 0 at signal GNT_MOT_PWR_EN and return result to PC.
| 8.9.27 | Connect Scope CH3 to 96V_24V_GNT current sample, CH2 to TP118 (OC_V_GNT_MOT_DRV). | DC Power Supply Resistive Load Oscilloscope | N/A | NO | N/A
| 8.9.28 | Send command to FPGA to set PWM duty cycle to 96%. | DC Power Supply Resistive Load Oscilloscope | PWM duty cycle set to 96%. | Yes | Set PWM duty cycle to 96% and return result to PC.
| 8.9.29 | Adjust resistive loads to 0.47 Ohms delta connection to GNT_DRV_PHASE_A, B, and C. (To get 96V_24V_GNT_DRV current ~70A) | DC Power Supply Resistive Load Oscilloscope | N/A | NO | N/A
| 8.9.30 | Send command to FPGA to output 100ms pulse at signal GNT_MOT_PWR_EN to generate a brief high current pulse for 100ms. | DC Power Supply Resistive Load Oscilloscope | GNT_MOT_PWR_EN get 100ms pulse. | Yes | Output 100ms pulse at signal GNT_MOT_PWR_EN to generate a brief high current pulse for 100ms and return result to PC.
| 8.9.31 | Send powerIfRd command to FPGA to read and verify the status of signal GNT_MOT_PWR_FLT# is HIGH. | DC Power Supply Resistive Load Oscilloscope | GNT_MOT_PWR_FLT# is high. | Yes | Read the status of signal GNT_MOT_PWR_FLT# and return result to PC.
| 8.9.32 | Scope will measure the peak current 62.9A±5% on CH3 and positive pulse of 4uS±5% on CH2. | DC Power Supply Resistive Load Oscilloscope | 1) CH3 peak current is 62.9A±5%. 2) CH2 positive pulse of 4uS±5%. | NO | N/A
| 8.9.33 | Disconnect Scope CH3 and CH2. Send command to FPGA to set PWM duty cycle back as before. | DC Power Supply Resistive Load | GNT_PWM_PHA_HI/LO, GNT_PWM_PHB_HI/LO and GNT_PWM_PHC_HI/LO get 333.3Hz pulses. | Yes | Output 333.3Hz pulses to signals GNT_PWM_PHA_HI/LO, GNT_PWM_PHB_HI/LO and GNT_PWM_PHC_HI/LO and return result to PC.
| 8.9.34 | Recover the PWM pulse input like 8.9.17 and adjust the resistive loads back to 25 ohms delta connection to GNT_DRV_PHASE_A, B, and C. | DC Power Supply Resistive Load | GNT_PWM_PHA_HI/LO, GNT_PWM_PHB_HI/LO and GNT_PWM_PHC_HI/LO get 333.3Hz pulses. | Yes | Output 333.3Hz pulses to signals GNT_PWM_PHA_HI/LO, GNT_PWM_PHB_HI/LO and GNT_PWM_PHC_HI/LO and return result to PC.
| 8.9.35 | Send gantMotGpio command to FPGA to output logic 1 to signal GNT_MOT_PWR_EN | DC Power Supply Resistive Load | GNT_MOT_PWR_EN set to 1. | Yes | Output logic 1 at signal GNT_MOT_PWR_EN and return result to PC.
| 8.9.36 | Send powerIfRd command to FPGA to read and very the status of signal OC_V_GNT_MOT_DRV is low. | DC Power Supply Resistive Load | OC_V_GNT_MOT_DRV is low. | Yes | Read the status of signal OC_V_GNT_MOT_DRV and return result to PC.
| 8.9.37 | Send gantMotGpio command to FPGA to output logic 0 to signal GNT_PWM_EN. | DC Power Supply Resistive Load | GNT_PWM_EN set to 0. | Yes | Output logic 0 at signal GNT_PWM_EN and return result to PC.
| 8.9.38 | Disconnect the loads from GNT_DRV_PHASE_A, B, and C. | DC Power Supply | N/A | NO | N/A
| 8.9.39 | Send gantry96VoltsEn command to FPGA to output logic 0 to signal 24V_GNT_EMOPS_EN. | DC Power Supply | 24V_GNT_EMOPS_EN set to 0. | Yes | Output logic 0 at signal 24V_GNT_EMOPS_EN and return result to PC.
| 8.9.40 | Send powerIfRd command to FPGA to read and verify the status of signal 24V_GNT_EMOPS_PG is low. | DC Power Supply | 24V_GNT_EMOPS_PG is low. | Yes | Read the status of signal 24V_GNT_EMOPS_PG and return result to PC.
| 8.9.41 | Remove 24V source from 24V_IN. | DC Power Supply | N/A | NO | N/A
| 8.9.42 | Apply 96VDC between 96V_GNT(J17.2,J17.3) and MGND. | DC Power Supply | N/A | NO | N/A
| 8.9.43 | Send gantry96VoltsEn command to FPGA to output logic 0 to signal GNT_EMOPS_EN | DC Power Supply | GNT_EMOPS_EN set to 0. | Yes | Output logic 0 at signal GNT_EMOPS_EN and return result to PC.
| 8.9.44 | Verify the LED D73 is ON. | DC Power Supply FEASA | LED D73 is on. | NO | N/A
| 8.9.45 | Select signal 24V_IN_MON to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage of signal 24V_IN(J18.1,J18.4) to be 0V±500mV. | DC Power Supply | Voltage at 24V_IN is 0V±500mV. | Yes | Select the signal 24V_IN_MON to be an input of the ADC, query the voltage at signal 24V_IN and return result to PC.
| 8.9.46 | Select signal 24V_GNT_EMOPS_S_MON to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage of signal 24V_GNT_EMOPS_S to be 0V±500mV. | DC Power Supply | Voltage at 24V_GNT_EMOPS_S is 0V±500mV. | Yes | Select the signal 24V_GNT_EMOPS_S_MON to be an input of the ADC, query the voltage at signal 24V_GNT_EMOPS_S and return result to PC.
| 8.9.47 | Select signal 96V_24V_GNT_MON to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage of signal 96V_24V_GNT(TP106) to be 96V±10%. | DC Power Supply | Voltage at 96V_24V_GNT is 96V±10%. | Yes | Select the signal 96V_24V_GNT_MON to be an input of the ADC, query the voltage at signal 96V_24V_GNT and return result to PC.
| 8.9.48 | Select signal 96V_GNT_MON to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage of signal 96V_GNT(J17.3,J17.4) to be 96V±10%. | DC Power Supply | Voltage at 96V_GNT is 96V±10%. | Yes | Select the signal 96V_GNT_MON to be an input of the ADC, query the voltage at signal 96V_GNT and return result to PC.
| 8.9.49 | Send gantMotEn and gantMotPwmSel command to FPGA to output 333.3Hz pulses to GNT_PWM_PHA_HI/LO, GNT_PWM_PHB_HI/LO and GNT_PWM_PHC_HI/LO. | DC Power Supply | GNT_PWM_PHA_HI/LO, GNT_PWM_PHB_HI/LO and GNT_PWM_PHC_HI/LO get 333.3Hz pulses. | Yes | Output 333.3Hz pulses to signals GNT_PWM_PHA_HI/LO, GNT_PWM_PHB_HI/LO and GNT_PWM_PHC_HI/LO and return result to PC.
| 8.9.50 | Send gantMotGpio command to FPGA to output logic 1 to signal GNT_PWM_EN | DC Power Supply | GNT_PWM_EN set to 1. | Yes | Output logic 1 at signal GNT_PWM_EN and return result to PC.
| 8.9.51 | Measure the signal amplitude of GNT_DRV_PHASE_A, B, and C to be 96V±10%. | DC Power Supply Oscilloscope | Amplitude of GNT_DRV_PHASE_A, B, and C is 96V±10%. | NO | N/A
| 8.9.52 | Measure the signal frequency of GNT_DRV_PHASE_A, B, and C to be 333.3HZ±5%. | DC Power Supply Oscilloscope | Frequency of GNT_DRV_PHASE_A, B, and C is 333.3Hz±5%. | NO | N/A
| 8.9.53 | Measure the phase of GNT_DRV_PHASE_A, B, and C to be 115 degrees to 125 degrees. | DC Power Supply Oscilloscope | Phase between GNT_DRV_PHASE_A, B, and C is within 115 and 125 degrees. | NO | N/A
| 8.9.54 | Connect 100Ohms resistive loads WYE connection to GNT_DRV_PHASE_A, B, and C. | DC Power Supply Resistive Load | N/A | NO | N/A
| 8.9.55 | Select 96V_24V_GNT_IMON_V to be input of ADC. Send adcAcquire command to FPGA to query and verify the voltage at signal Select 96V_24V_GNT_IMON_V to be within 25mV and 50mV. (Spec limit to be settled with KGB.) | DC Power Supply Resistive Load | Voltage at 96V_24V_GNT_IMON_V is 25mV ~ 50mV. | Yes | Select the signal 96V_24V_GNT_IMON_V to be an input of the ADC, query the voltage at signal 96V_24V_GNT_IMON_V and return result to PC.
| 8.9.56 | Select signal GNT_DRV_CURA to be the input of the ADC. Send adc1024Samples command to FPGA to read and verify the voltage at signal GNT_DRV_CURA(TP91) to be between 20mV to 76.6mV. (Spec limit to be settled with KGB.) | DC Power Supply Resistive Load | Voltage at GNT_DRV_CURA is 20mV ~ 76.6mV. | Yes | Select the signal GNT_DRV_CURA to be an input of the ADC, query the voltage at signal GNT_DRV_CURA and return result to PC.
| 8.9.57 | Select signal GNT_DRV_CURB to be the input of the ADC. Send adc1024Samples command to FPGA to read and verify the voltage at signal GNT_DRV_CURB(TP69) to be between 20mV to 76.6mV. (Spec limit to be settled with KGB.) | DC Power Supply Resistive Load | Voltage at GNT_DRV_CURB is 20mV ~ 76.6mV. | Yes | Select the signal GNT_DRV_CURB to be an input of the ADC, query the voltage at signal GNT_DRV_CURB and return result to PC.
| 8.9.58 | Disable 96V_GNT, remove the loads from J22 | DC Power Supply | N/A | NO | N/A
| 8.9.59 | Connect 5ohms between J13.1 (GNT_BRK1) and J13.10 (GNT_BRK1_RET), J13.5 (GNT_BRK2) and J13.12 (GNT_BRK2_RET), J13.7 (GNT_BRK3) and J13.14 (GNT_BRK3_RET). | DC Power Supply Resistive Load | N/A | NO | N/A
| 8.9.60 | Send gantry96VoltsEn command to FPGA to output logic 1 to signal GNT_BRK_PWR_EN. | DC Power Supply Resistive Load | GNT_BRK_PWR_EN set to 1. | Yes | Output logic 1 at signal GNT_BRK_PWR_EN and return result to PC.
| 8.9.61 | Send command to FPGA to output 5s pulses to GNT_BRK1_PWM_HI/LO, GNT_BRK1_RET_PWM_HI/LO; GNT_BRK2_PWM_HI/LO, GNT_BRK2_RET_PWM_HI/LO; GNT_BRK3_PWM_HI/LO, GNT_BRK3_RET_PWM_HI/LO. | DC Power Supply Resistive Load | GNT_BRK1_PWM_HI/LO, GNT_BRK1_RET_PWM_HI/LO; GNT_BRK2_PWM_HI/LO, GNT_BRK2_RET_PWM_HI/LO; GNT_BRK3_PWM_HI/LO, GNT_BRK3_RET_PWM_HI/LO get 5s pulse | Yes | Output 5s pulses to GNT_BRK1_PWM_HI/LO, GNT_BRK1_RET_PWM_HI/LO; GNT_BRK2_PWM_HI/LO, GNT_BRK2_RET_PWM_HI/LO; GNT_BRK3_PWM_HI/LO, GNT_BRK3_RET_PWM_HI/LO and return result to PC.
| 8.9.62 | Select signal GNT_BRK1_CUR to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage to be 1.7V±10%. (Spec limit TBD.) | DC Power Supply Resistive Load | Voltage at GNT_BRK1_CUR is 1.7V±10%. | Yes | Select the signal GNT_BRK1_CUR to be an input of the ADC, query the voltage at signal GNT_BRK1_CUR and return result to PC.
| 8.9.63 | Select signal GNT_BRK2_CUR to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage to be 1.7V±10%. (Spec limit TBD.) | DC Power Supply Resistive Load | Voltage at GNT_BRK2_CUR is 1.7V±10%. | Yes | Select the signal GNT_BRK2_CUR to be an input of the ADC, query the voltage at signal GNT_BRK2_CUR and return result to PC.
| 8.9.64 | Select signal GNT_BRK3_CUR to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage to be 1.7V±10%. (Spec limit TBD.) | DC Power Supply Resistive Load | Voltage at GNT_BRK3_CUR is 1.7V±10%. | Yes | Select the signal GNT_BRK3_CUR to be an input of the ADC, query the voltage at signal GNT_BRK3_CUR and return result to PC.
| 8.9.65 | Send command to FPGA to output 0.5s pulses to above GNT_BRK_PWM location. | DC Power Supply Resistive Load | GNT_BRK1_PWM_HI/LO, GNT_BRK1_RET_PWM_HI/LO; GNT_BRK2_PWM_HI/LO, GNT_BRK2_RET_PWM_HI/LO; GNT_BRK3_PWM_HI/LO, GNT_BRK3_RET_PWM_HI/LO get 0.5s pulse | Yes | Output 0.5s pulses to GNT_BRK1_PWM_HI/LO, GNT_BRK1_RET_PWM_HI/LO; GNT_BRK2_PWM_HI/LO, GNT_BRK2_RET_PWM_HI/LO; GNT_BRK3_PWM_HI/LO, GNT_BRK3_RET_PWM_HI/LO and return result to PC.
| 8.9.66 | Select signal GNT_BRK1_CUR to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage to be 3.5V±10%. (Spec limit TBD.) | DC Power Supply Resistive Load | Voltage at GNT_BRK1_CUR is 3.5V±10%. | Yes | Select the signal GNT_BRK1_CUR to be an input of the ADC, query the voltage at signal GNT_BRK1_CUR and return result to PC.
| 8.9.67 | Select signal GNT_BRK2_CUR to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage to be 3.5V±10%. (Spec limit TBD.) | DC Power Supply Resistive Load | Voltage at GNT_BRK2_CUR is 3.5V±10%. | Yes | Select the signal GNT_BRK2_CUR to be an input of the ADC, query the voltage at signal GNT_BRK2_CUR and return result to PC.
| 8.9.68 | Select signal GNT_BRK3_CUR to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage to be 3.5V±10%. (Spec limit TBD.) | DC Power Supply Resistive Load | Voltage at GNT_BRK3_CUR is 3.5V±10%. | Yes | Select the signal GNT_BRK3_CUR to be an input of the ADC, query the voltage at signal GNT_BRK3_CUR and return result to PC.
| 8.9.69 | Send command to FPGA to disable pulses from above GNT_BRK_PWM location. Remove 5ohm from J13. | DC Power Supply | GNT_BRK1_PWM_HI/LO, GNT_BRK1_RET_PWM_HI/LO; GNT_BRK2_PWM_HI/LO, GNT_BRK2_RET_PWM_HI/LO; GNT_BRK3_PWM_HI/LO, GNT_BRK3_RET_PWM_HI/LO pulse disabled | Yes | Disable pulses to GNT_BRK1_PWM_HI/LO, GNT_BRK1_RET_PWM_HI/LO; GNT_BRK2_PWM_HI/LO, GNT_BRK2_RET_PWM_HI/LO; GNT_BRK3_PWM_HI/LO, GNT_BRK3_RET_PWM_HI/LO and return result to PC.
| 8.9.70 | Send motgpoWr command to FPGA to output logic 1 to signal GNT_BRK_SW_MON. | DC Power Supply | GNT_BRK_SW_MON set to 1. | Yes | Output logic 1 at signal GNT_BRK_SW_MON and return result to PC.
| 8.9.71 | Measure J13.2, J13.4 and J13.6 (GNT_BRK_SW_24V) to be between 23.5V and 24.5V. | DC Power Supply DMM | Voltage at J13.2, J13.4 and J13.6 is 24V±0.5V. | NO | N/A
| 8.9.72 | Send motgpoWr command to FPGA to output logic 0 to signal GNT_BRK_SW_MON. | DC Power Supply | GNT_BRK_SW_MON set to 0. | Yes | Output logic 0 at signal GNT_BRK_SW_MON and return result to PC.
| 8.9.73 | Measure J13.2, J13.4 and J13.6 (GNT_BRK_SW_24V) to be 0V±0.5V. | DC Power Supply DMM | Voltage at J13.2, J13.4 and J13.6 is 0V±0.5V. | NO | N/A
| 8.9.74 | Apply 24V to J13.11 (GNT_BRK1_SW_FB), J13.13 (GNT_BRK2_SW_FB) and J13.15 (GNT_BRK3_SW_FB). | DC Power Supply | N/A | NO | N/A
| 8.9.75 | Send command to FPGA to verify the status of signals GNT_BRK1_FB#, GNT_BRK2_FB# and GNT_BRK3_FB# to be logic low. | DC Power Supply | GNT_BRK1_FB#, GNT_BRK2_FB# and GNT_BRK3_FB# is low. | Yes | Read the status of signal GNT_BRK1_FB#, GNT_BRK2_FB# and GNT_BRK3_FB# and return result to PC.
| 8.9.76 | Check D84, D85 and D86 are illuminated. | DC Power Supply FEASA | D84, D85 and D86 are on. | NO | N/A
| 8.9.77 | Apply 0V to J13.11 (GNT_BRK1_SW_FB), J13.13 (GNT_BRK2_SW_FB) and J13.15 (GNT_BRK3_SW_FB). | DC Power Supply | N/A | NO | N/A
| 8.9.78 | Send command to FPGA to verify the status of signals GNT_BRK1_FB#, GNT_BRK2_FB# and GNT_BRK3_FB# to be logic high. | DC Power Supply | GNT_BRK1_FB#, GNT_BRK2_FB# and GNT_BRK3_FB# is high. | Yes | Read the status of signal GNT_BRK1_FB#, GNT_BRK2_FB# and GNT_BRK3_FB# and return result to PC.
| 8.9.79 | Check D84, D85 and D86 are not illuminated. | DC Power Supply FEASA | D84, D85 and D86 are off. | NO | N/A
| 8.9.80 | Send gantry96VoltsEn command to FPGA to output logic 0 to signal GNT_BRK_PWR_EN. | DC Power Supply | GNT_BRK_PWR_EN set to 0. | Yes | Output logic 0 at signal GNT_BRK_PWR_EN and return result to PC.
| 8.9.81 | Connect Scope CH3 to 96V_GNT_BRK_DRV current sample, CH2 to TP39 (OC_V_GNT_BRK_DRV). | DC Power Supply Oscilloscope | N/A | NO | N/A
| 8.9.82 | Send command to FPGA to set PWM duty cycle to 96% and apply PWM signal to GNT_BRK1_PWM_HI/LO, GNT_BRK2_PWM_HI/LO and GNT_BRK3_PWM_HI/LO. | DC Power Supply | GNT_BRK1_PWM_HI/LO, GNT_BRK2_PWM_HI/LO and GNT_BRK3_PWM_HI/LO applied with 96% duty cycle PWM | Yes | Set PWM duty cycle to 96% and apply PWM signal to GNT_BRK1_PWM_HI/LO, GNT_BRK2_PWM_HI/LO and GNT_BRK3_PWM_HI/LO and return result to PC.
| 8.9.83 | Adjust resistive loads to 0.98 ohms delta connection to J13.1 (GNT_BRK1), J13.5 (GNT_BRK2) and J13.7 (GNT_BRK3). (To get 96V_GNT_BRK_DRV current ~34A) | DC Power Supply Resistive Load | N/A | NO | N/A
| 8.9.84 | Apply 24VDC between signal 96V_GNT and MGND. Send command to FPGA to output 100ms pulse at signal GNT_BRK_PWR_EN to generate a brief high current pulse for 100ms. | DC Power Supply Resistive Load | GNT_BRK_PWR_EN get 100ms pulse | Yes | Output 100ms pulse at signal GNT_BRK_PWR_EN to generate a brief high current pulse for 100ms and return result to PC.
| 8.9.85 | Send powerIfRd command to FPGA to read and verify the status of signal GNT_BRK_PWR_FLT# is HIGH. | DC Power Supply Resistive Load | GNT_BRK_PWR_FLT# is high. | Yes | Read the status of signal GNT_BRK_PWR_FLT# and return result to PC.
| 8.9.86 | Scope will measure the peak current 30.26A±5% on CH3 and positive pulse of 4uS±5% on CH2. | DC Power Supply Oscilloscope Resistive Load | 1) CH3 peak current is 30.26A±5%. 2) CH2 positive pulse of 4uS±5%. | NO | N/A
| 8.9.87 | Disconnect Scope CH3 and CH2. Remove resistive load from J13. Send command to FPGA to disable PWM signals. | DC Power Supply | GNT_BRK1_PWM_HI/LO, GNT_BRK2_PWM_HI/LO and GNT_BRK3_PWM_HI/LO PWM disabled | Yes | Disable pulses to GNT_BRK1_PWM_HI/LO, GNT_BRK2_PWM_HI/LO and GNT_BRK3_PWM_HI/LO and return result to PC.
| 8.9.88 | Send gantMotGpio command to FPGA to output logic 0 to signal GNT_PWM_EN. | DC Power Supply | GNT_PWM_EN set to 0. | Yes | Output logic 0 at signal GNT_PWM_EN and return result to PC.
| 8.9.89 | Send gantry96VoltsEn command to FPGA to output logic 0 to signal GNT_MOT_PWR_EN. | DC Power Supply | GNT_MOT_PWR_EN set to 0. | Yes | Output logic 0 at signal GNT_MOT_PWR_EN and return result to PC.
8.10 Lift 96V Motor Test | 8.10.1 | Apply +24VDC to 24V_IN. Connect J14.2 (LFT_96VSHUNT-) to scope CH3. | DC Power Supply Oscilloscope | N/A | NO | N/A
| 8.10.2 | Apply 96VDC between signal 96V_LFT and MGND. Connect 50ohms resistor across J14.1 and J14.2. | DC Power Supply Oscilloscope Resistive Load | N/A | NO | N/A
| 8.10.3 | Send lift96VoltsEn command to FPGA to output logic 0 to signal LFT_EMOPS_EN and logic 1 to LFT_MOT_PWR_EN, measure the voltage at J14.1 and GND to be 96V±5%. | DC Power Supply DMM Resistive Load | 1) LFT_EMOPS_EN set to 0 2) LFT_MOT_PWR_EN set to 1 3) Voltage at J14.1 is 96V±5%. | Yes | Output logic 0 at signal LFT_EMOPS_EN and logic 1 to LFT_MOT_PWR_EN and return result to PC.
| 8.10.4 | Send command to FPGA to output logic 1 to LFT_SHUNT_EN_OUT for 2.5s | DC Power Supply Resistive Load | LFT_SHUNT_EN_OUT set to 1 for 2.5s. | Yes | Output logic 1 to signal LFT_SHUNT_EN_OUT for 2.5s and return result to PC.
| 8.10.5 | Measure the voltage at J14.2 (LFT_96VSHUNT-) from Scope CH3, it should be a 96V-0V pulse for 2.5s period. | DC Power Supply Oscilloscope Resistive Load | Voltage at J14.2 is 96V-0V pulse for 2.5s period. | NO | N/A
| 8.10.6 | Disconnect 50ohms resistor from J14.1, disconnect CH3 from J14.2. | DC Power Supply | N/A | NO | N/A
| 8.10.7 | Send lift96VoltsEn command to FPGA to output logic 1 to signal LFT_EMOPS_EN, and LFT_SHUNT_EN_OUT. | DC Power Supply | 1) LFT_EMOPS_EN set to 1. 2) LFT_SHUNT_EN_OUT set to 1. | Yes | Output logic 1 to signal LFT_EMOPS_EN, and LFT_SHUNT_EN_OUT and return result to PC.
| 8.10.8 | Apply 24VDC to signal 24V_IN and GND. | DC Power Supply | N/A | NO | N/A
| 8.10.9 | Send powerIfRd command to FPGA to read and verify the status of signal 24V_LFT_EMOPS_PG is low. | DC Power Supply | 24V_LFT_EMOPS_PG is low. | Yes | Read the status of signal 24V_LFT_EMOPS_PG and return result to PC.
| 8.10.10 | Send lift96VoltsEn to FPGA to output logic 1 to 24V_LFT_EMOPS_EN. | DC Power Supply | 24V_LFT_EMOPS_EN set to 1. | Yes | Output logic 1 to signal 24V_LFT_EMOPS_EN and return result to PC.
| 8.10.11 | Send powerIfRd command to FPGA to read and verify the status of signal 24V_LFT_EMOPS_PG is high. | DC Power Supply | 24V_LFT_EMOPS_PG is high. | Yes | Read the status of signal 24V_LFT_EMOPS_PG and return result to PC.
| 8.10.12 | Select signal 24V_LFT_EMOPS_S_MON to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage at signal 24V_LFT_EMOPS_S to be 24V±5%. | DC Power Supply | Voltage at 24V_LFT_EMOPS_S is 21.5V~ 24.5V. | Yes | Select the signal 24V_LFT_EMOPS_S_MON to be an input of the ADC, query the voltage at signal 24V_LFT_EMOPS_S and return result to PC.
| 8.10.13 | Select signal 96V_24V_LFT_MON to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage at signal 96V_24V_LFT to be between 21.5V and 24.5V. | DC Power Supply | Voltage at 96V_24V_LFT is 21.5V~ 24.5V. | Yes | Select the signal 96V_24V_LFT_MON to be an input of the ADC, query the voltage at signal 96V_24V_LFT and return result to PC.
| 8.10.14 | Verify the output voltage of U8 is between 0.6V and 0.96V. | DC Power Supply | Output voltage of U8 is between 0.6V and 0.96V. | Yes | Read the output voltage of U8 and return result to PC.
| 8.10.15 | Send lifttMotPwmSel and liftMotEn command to PFGA to output 333.3Hz pulses to signals LFT_PWM_PHA_HI/LO, LFT_PWM_PHB_HI/LO and LFT_PWM_PHC_HI/LO. | DC Power Supply | LFT_PWM_PHA_HI/LO, LFT_PWM_PHB_HI/LO and LFT_PWM_PHC_HI/LO applied with 333.3Hz pulses | Yes | Output 333.3Hz pulses to signals LFT_PWM_PHA_HI/LO, LFT_PWM_PHB_HI/LO and LFT_PWM_PHC_HI/LO.
| 8.10.16 | Send liftMotGpio command to FPGA to output logic 1 at signals LFT_PWM_EN. | DC Power Supply | LFT_PWM_EN set to 1. | Yes | Output logic 1 to signal LFT_PWM_EN and return result to PC.
| 8.10.17 | Measure the signal amplitude of LFT_DRV_PHASE_A, B, and C to be 24V±5%. | DC Power Supply Oscilloscope | Amplitude of LFT_DRV_PHASE_A, B, and C is 24V±5%. | NO | N/A
| 8.10.18 | Measure the signal frequency of LFT_DRV_PHASE_A, B, and C to be 333.3Hz±5%. | DC Power Supply Oscilloscope | Frequency of LFT_DRV_PHASE_A, B, and C is 333.3Hz±5%. | NO | N/A
| 8.10.19 | Measure the phase between LFT_DRV_PHASE_A, B, and C to be within 115 and 125 degrees. | DC Power Supply Oscilloscope | Phase between LFT_DRV_PHASE_A, B, and C is within 115 and 125 degrees. | NO | N/A
| 8.10.20 | Send liftMotGpio command to FPGA to output logic 0 to signals LFT_PWM_EN. | DC Power Supply | LFT_PWM_EN set to 0. | Yes | Output logic 0 to signal LFT_PWM_EN and return result to PC.
| 8.10.21 | Connect 25Ohms resistive loads delta connection to LFT_DRV_PHASE_A, B, and C. | DC Power Supply Resistive Load | N/A | NO | N/A
| 8.10.22 | Send liftMotGpio command to FPGA to output logic 1 to signals LFT_PWM_EN. | DC Power Supply Resistive Load | LFT_PWM_EN set to 1. | Yes | Output logic 1 to signal LFT_PWM_EN and return result to PC.
| 8.10.23 | Select the signal 96V_24V_LFT_IMON_V to be an input of the ADC. Send adcAcquire command to FPGA to query and verify the voltage at signal 96V_24V_LFT_IMON_V to be within 87.5 mV and 112.5mV. (Spec limit to be settled with KGB.) | DC Power Supply Resistive Load | Voltage at 96V_24V_LFT_IMON_V is 87.5 mV ~ 112.5mV. | Yes | Select the signal 96V_24V_LFT_IMON_V to be an input of the ADC, query the voltage at signal 96V_24V_LFT_IMON_V and return result to PC.
| 8.10.24 | Send gantry96VoltsEn command to FPGA to output logic 0 to signal LFT_MOT_PWR_EN. | DC Power Supply Resistive Load | LFT_MOT_PWR_EN set to 0. | Yes | Output logic 0 to signal LFT_MOT_PWR_EN and return result to PC.
| 8.10.25 | Connect Scope CH3 to 96V_24V_LFT current sample, CH2 to TP38 (OC_V_LFT_MOT_DRV). | DC Power Supply Oscilloscope Resistive Load | N/A | NO | N/A
| 8.10.26 | Send command to FPGA to set PWM duty cycle to 96%. | DC Power Supply Resistive Load | LFT_PWM_PHA_HI/LO, LFT_PWM_PHB_HI/LO and LFT_PWM_PHC_HI/LO applied with 96% duty cycle PWM | Yes | Set PWM duty cycle to 96% and apply PWM signal to LFT_PWM_PHA_HI/LO, LFT_PWM_PHB_HI/LO and LFT_PWM_PHC_HI/LO and return result to PC.
| 8.10.27 | Adjust resistive loads to 1.23Ohms delta connection to LFT_DRV_PHASE_A, B, and C. (To get 96V_24V_LFT_DRV current ~27A) | DC Power Supply Resistive Load | N/A | NO | N/A
| 8.10.28 | Send command to FPGA to output 100ms pulse at signal LFT_MOT_PWR_EN to generate a brief high current pulse for 100ms. | DC Power Supply Resistive Load | LFT_MOT_PWR_EN get 100ms pulse | Yes | Output 100ms pulse at signal LFT_MOT_PWR_EN to generate a brief high current pulse for 100ms and return result to PC.
| 8.10.29 | Send powerIfRd command to FPGA to read and verify the status of signal LFT_MOT_PWR_FLT# is HIGH. | DC Power Supply Resistive Load | LFT_MOT_PWR_FLT# is high. | Yes | Read the status of signal LFT_MOT_PWR_FLT# and return result to PC.
| 8.10.30 | Scope will measure the peak current 24.608A±5% on CH3 and positive pulse of 4uS±5% on CH2. | DC Power Supply Oscilloscope Resistive Load | 1) CH3 peak current is 24.608A±5%. 2) CH2 positive pulse of 4uS±5%. | NO | N/A
| 8.10.31 | Disconnect Scope CH3 and CH2. Send command to FPGA to set PWM duty cycle back as before. | DC Power Supply Resistive Load | LFT_PWM_PHA_HI/LO, LFT_PWM_PHB_HI/LO and LFT_PWM_PHC_HI/LO applied with 333.3Hz pulses | Yes | Output 333.3Hz pulses to signals LFT_PWM_PHA_HI/LO, LFT_PWM_PHB_HI/LO and LFT_PWM_PHC_HI/LO.
| 8.10.32 | Recover the PWM pulse input like 8.10.15, adjust resistive loads back to 25 Ohm delta connection to LFT_DRV_PHASE_A, B, and C. | DC Power Supply Resistive Load | LFT_PWM_PHA_HI/LO, LFT_PWM_PHB_HI/LO and LFT_PWM_PHC_HI/LO applied with 333.3Hz pulses | Yes | Output 333.3Hz pulses to signals LFT_PWM_PHA_HI/LO, LFT_PWM_PHB_HI/LO and LFT_PWM_PHC_HI/LO.
| 8.10.33 | Send gantMotGpio command to FPGA to output logic 1 to signal LFT_MOT_PWR_EN. | DC Power Supply Resistive Load | LFT_MOT_PWR_EN set to 1. | Yes | Output logic 1 to signal LFT_MOT_PWR_EN and return result to PC.
| 8.10.34 | Send powerIfRd command to FPGA to read and very the status of signal LFT_OVER_CURR is low. | DC Power Supply Resistive Load | LFT_OVER_CURR is low. | Yes | Read the status of signal LFT_OVER_CURR and return result to PC.
| 8.10.35 | Send liftMotGpio command to FPGA to output logic 0 to signal LFT_PWM_EN. | DC Power Supply Resistive Load | LFT_PWR_EN set to 0. | Yes | Output logic 0 to signal LFT_PWM_EN and return result to PC.
| 8.10.36 | Disconnect the loads from LFT_DRV_PHASE_A, B, and C. | DC Power Supply | N/A | NO | N/A
| 8.10.37 | Send lift96VoltsEn command to FPGA to output logic 0 to signal 24V_LFT_EMOPS_EN. | DC Power Supply | 24V_LFT_EMOPS_EN set to 0. | Yes | Output logic 0 to signal 24V_LFT_EMOPS_EN and return result to PC.
| 8.10.38 | Send powerIfRd command to FPGA to read and verify the status of signal 24V_LFT_EMOPS_PG is low. | DC Power Supply | 24V_LFT_EMOPS_PG is low. | Yes | Read the status of signal 24V_LFT_EMOPS_PG and return result to PC.
| 8.10.39 | Remove 24V from 24V_IN. | DC Power Supply | N/A | NO | N/A
| 8.10.40 | Apply 96VDC between 96V_LFT and MGND. | DC Power Supply | N/A | NO | N/A
| 8.10.41 | Send lift96VoltsEn command to FPGA to output logic 0 to signal LFT_EMOPS_EN. | DC Power Supply | LFT_EMOPS_EN set to 0. | Yes | Output logic 0 to signal LFT_EMOPS_EN and return result to PC.
| 8.10.42 | Verify the LED D36 is ON. | DC Power Supply FEASA | D36 is on | NO | N/A
| 8.10.43 | Select signal 24V_LFT_EMOPS_S_MON to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage of signal 24V_LFT_EMOPS_S to be 0V±500mV. | DC Power Supply | Voltage at 24V_LFT_EMOPS_S is 0V±500mV. | Yes | Select the signal 24V_LFT_EMOPS_S_MON to be an input of the ADC, query the voltage at signal 24V_LFT_EMOPS_S and return result to PC.
| 8.10.44 | Select signal 96V_24V_LFT_MON to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage of signal 96V_24V_LFT to be 96V±10%. | DC Power Supply | Voltage at 96V_24V_LFT is 96V±10%. | Yes | Select the signal 96V_24V_LFT_MON to be an input of the ADC, query the voltage at signal 96V_24V_LFT and return result to PC.
| 8.10.45 | Select signal 96V_LFT_MON to be the input of ADC. Send adcAcquire command to FPGA to query and verify the voltage of signal 96V_LFT to be 96V±10%. | DC Power Supply | Voltage at 96V_LFT is 96V±10%. | Yes | Select the signal 24V_LFT_EMOPS_S_MON to be an input of the ADC, query the voltage at signal 24V_LFT_EMOPS_S and return result to PC.
| 8.10.46 | Send liftMotEn and liftMotPwmSel command to FPGA to output 333.3Hz pulses to LFT_DRV_PHASE_A, B, and C. | DC Power Supply | LFT_DRV_PHASE_A, B, and C applied with 333.3Hz pulses | Yes | Output 333.3Hz pulses to signals LFT_DRV_PHASE_A, B, and C.
| 8.10.47 | Send liftMotGpio command to FPGA to output logic 1 to signal LFT_PWM_EN. | DC Power Supply | LFT_PWM_EN set to 1. | Yes | Output logic 1 to signal LFT_PWM_EN and return result to PC.
| 8.10.48 | Measure the signal amplitude of LFT_DRV_PHASE_A, B, and C to be 96V±5%. | DC Power Supply Oscilloscope | Amplitude of LFT_DRV_PHASE_A, B, and C is 96V±10%. | NO | N/A
| 8.10.49 | Measure the signal frequency of LFT_DRV_PHASE_A, B, and C to be 333.3HZ±5%. | DC Power Supply Oscilloscope | Frequency of LFT_DRV_PHASE_A, B, and C is 333.3Hz±5%. | NO | N/A
| 8.10.50 | Measure the phase of LFT_DRV_PHASE_A, B, and C to be 115 degrees to 125 degrees. | DC Power Supply Oscilloscope | Phase between LFT_DRV_PHASE_A, B, and C is within 115 and 125 degrees. | NO | N/A
| 8.10.51 | Connect 100Ohms resistive loads WYE connection to LFT_DRV_PHASE_A, B, and C. | DC Power Supply Resistive Load | N/A | NO | N/A
| 8.10.52 | Select 96V_24V_LFT_IMON_V to be input of ADC. Send adcAcquire command to FPGA to query and verify the voltage at signal 96V_24V_LFT_IMON_V to be within 25mV and 50mV. (Spec limit to be settled with KGB.) | DC Power Supply Resistive Load | Voltage at 24V_LFT_IMON_V is 25mV ~ 50mV. | Yes | Select the signal 96V_24V_LFT_IMON_V to be an input of the ADC, query the voltage at signal 24V_LFT_IMON_V and return result to PC.
| 8.10.53 | Select signal LFT_DRV_CURA to be the input of the ADC. Send adc1024Samples command to FPGA to read and verify the voltage at signal LFT_DRV_CURA to be between 20mV to 76.6mV. (Spec limit to be settled with KGB.) | DC Power Supply Resistive Load | Voltage at LFT_DRV_CURA is 20mV ~ 76.6mV. | Yes | Select the signal LFT_DRV_CURA to be an input of the ADC, query the voltage at signal LFT_DRV_CURA and return result to PC.
| 8.10.54 | Select signal LFT_DRV_CURB to be the input of the AD C. Send adc1024Samples command to FPGA to read and verify the voltage at signal LFT_DRV_CURB to be between 20mV to 76.6mV. (Spec limit to be settled with KGB.) | DC Power Supply Resistive Load | Voltage at LFT_DRV_CURB is 20mV ~ 76.6mV. | Yes | Select the signal LFT_DRV_CURB to be an input of the ADC, query the voltage at signal LFT_DRV_CURB and return result to PC.
| 8.10.55 | Disable 96V_LFT, remove the loads from J23. | DC Power Supply | N/A | NO | N/A
| 8.10.56 | Send liftMotGpio command to FPGA to output logic 0 to signal LFT_PWM_EN. | DC Power Supply | LFT_PWM_EN set to 0. | Yes | Output logic 0 to signal LFT_PWM_EN and return result to PC.
| 8.10.57 | Send gantry96VoltsEn command to FPGA to output logic 0 to signal LFT_MOT_PWR_EN. | DC Power Supply | LFT_MOT_PWR_EN set to 0. | Yes | Output logic 0 to signal LFT_MOT_PWR_EN and return result to PC.
8.11 Program | 8.11.1 | The test program will program the FPGA U54 with the customer provided firmware, Part#xxxxxxx(TBD).stp in the final if the board passes the test. After the programming sequence, the FPGA content has to be read back using “DEVICE_INFO” mode of FlashPro tool; compare the checksum with programming log (reference), saved in local drive. Checksums should match. | DC Power Supply Actel FlashPro4 programming emulator | FPGA U54 programmed with the customer provided firmware | NO | N/A
8.12 Power Down | 8.12.1 | Turns off the UUT power inputs. | N/A | N/A | NO | N/A
| 8.12.2 | Disconnect all cable assemblies from the UUT connectors. | N/A | N/A | NO | N/A
Note on B1: Delete Unused Rows